Consultation Questionnaire Exemption 2 of RoHS Annex IV

Current wording of the exemption:

Lead bearings in X-ray tubes

Expires in July 2021 for cat. 8 and 9 equipment other than in-vitro diagnostic devices and monitoring and control instruments in industry

# Acronyms and Definitions

# Background

Bio Innovation Service, UNITAR and Fraunhofer IZM have been appointed[[1]](#footnote-2) by the European Commission through for the evaluation of applications for the review of requests for new exemptions and the renewal of exemptions currently listed in Annexes III and IV of the RoHS Directive 2011/65/EU.

COCIR has submitted a request[[2]](#footnote-3) for the continuation of the above-mentioned exemption. The request has been subject to a first completeness and plausibility check. The applicant has been asked to answer additional questions and to provide additional information, available on the request webpage of the stakeholder consultation.[[3]](#footnote-4)

SUMMARY OF THE EXEMPTION REQUEST

The applicant requests the renewal of the exemption with the current wording and scope for the maximum possible 7 years.

According to the applicant2, “*This exemption is required for thin lead coatings on steel bearings used in rotating anode X-ray tubes. To avoid overheating of the focal track (portion of the anode where the electrons bombard) the target needs to rotate. The rotating anode is supported by a bearing, which has to be lubricated with vacuum compatible materials. Lead coated ball bearings are used due to their reliability and low noise. Low power X-ray tubes can use stationary targets (anodes) and some very high power designs used in CT equipment with high power X-ray tubes use liquid metal bearings, but neither are technically suitable in the types of X-ray equipment that currently use lead coated bearings.*"

The stakeholder consultation is part of the review process for the request at hand. The objective of this consultation and the review process is to collect and to evaluate information and evidence according to the criteria listed in Art. 5(1)(a) of Directive 2011/65/EU.[[4]](#footnote-5)

To contribute to this stakeholder consultation, please answer the questions below by December 2nd, 2020.

# Questions

1. The applicant has requested the renewal of exemption 2 of RoHS Annex IV with the current wording for seven years.
	1. Please let us know whether you support or disagree with the wording, scope and re-quested duration of the exemption. To support your views, please provide detailed technical argumentation / evidence in line with the criteria4 in Art. 5(1)(a).
	2. If applicable, please suggest an alternative wording and duration and explain your proposal.
2. Please provide information concerning possible substitutes or elimination possibilities at pre sent or in the future so that the requested exemption could be restricted or revoked.
	1. Please explain substitution and elimination possibilities and for which part of the ap-plications in the scope of the requested exemption they are relevant.
	2. Please provide information as to research to find alternatives that do not rely on the exemption under review (substitution or elimination), and which may cover part or all of the applications in the scope of the exemption request.
	3. Please provide a roadmap of such on-going substitution/elimination and research (phases that are to be carried out), detailing the current status as well as the estimated time needed for further stages.
3. Do you know of other manufacturers producing devices of comparable features and performance like the ones in the scope of this exemption request that do not depend on RoHS-restricted substances, or use smaller amounts of these substances compared to the applications in the scope of this requested exemption?
4. As part of the evaluation, socio-economic impacts shall also be compiled and evaluated. For this purpose, if you have information on socioeconomic aspects, please provide details in respect of the following:
	1. What are the volumes of EEE in the scope of the requested exemptions which are placed on the market per year?
	2. What are the volumes of additional waste to be generated should the requested ex-emption not be renewed or not be renewed for the requested duration?
	3. What are estimated impacts on employment in total, in the EU and outside the EU, should the requested exemption not be renewed or be renewed for less than the re-quested time period? Please detail the main sectors in which possible impacts are expected – manufacturers of equipment in the scope of the exemption, suppliers, re-tail, users of MRI devices, etc.
	4. Please estimate additional costs associated should the requested exemption not be renewed, and how this is divided between various sectors (e.g. private, public, industry: manufacturers, suppliers, retailers).
5. Is there any other information you wish to provide?

Please note that answers to these questions can be published in the stakeholder consultation, which is part of the evaluation of this request. If your answers contain confidential information, please provide a version that can be made public along with a confidential version, in which proprietary information is clearly marked.

Please do not forget to provide your contact details (Name, Organisation, e-mail and phone number) so that the project team can contact you in case there are questions concerning your contribution.

1. It is implemented through the specific contract 070201/2020/832829/ENV.B.3 under the Framework contract ENV.B.3/FRA/2019/0017 [↑](#footnote-ref-2)
2. Exemption request available at <https://rohs.biois.eu/Ex_2-IV_COCIR_Renewal-Request.pdf> [↑](#footnote-ref-3)
3. Clarification questionnaire available at <https://rohs.biois.eu/Ex_2-IV_COCIR_Questionnaire-1_Clarification.pdf> [↑](#footnote-ref-4)
4. Directive 2011/65/EU (RoHS) available at <http://eur-lex.europa.eu/LexUriServ/LexUriServ.do?uri=CELEX:32011L0065:EN:NOT> [↑](#footnote-ref-5)